## S&M 20 kHz data
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$S&M 20 kHz data

MP Interconsulting (Miodrag Prokic) - Carbagas bell horn project.
S&M FC2020 (2000 watt) power supply + CV15 0.8 mil (nominal) xdu + booster + 2" dia 1:1 horn
The horn is immersed approx 1/4" in tap water (not degassed).

	Face dia (inch)	Face area (in^2)	Xdu nominal ampl (microns)	Booster	Booster Gain (approx)	Relative variac amplitude	Approx horn amplitude with booster / (microns, p-p)	Approx horn amplitude with booster / (microns, peak)	Power in air (watts)	Power in water (watts)	Net cavitation power (watts)	Areal cavitation power density (watt/in^2)	Areal cavitation power density (watt/cm^2)	Areal cavitation power density (watt/mm^2)	Cavitation power density (watt/in^2/micron)	Cavitation power density (watt/cm^2/micron)
	2.000	3.142	20.3	Green	1.0	0.33	6.7	3.3	3.3	80	76.7	24	3.78	0.038	3.6	0.56
				Green	1.0	1.00	20.3	10.2	30	180	150	48	7.40	0.074	2.4	0.36
				Gold	1.5	1.00	30.5	15.2	52	300	248	79	12.24	0.122	2.6	0.40
				Silver	2.0	1.00	40.6	20.3	70	400	330	105	16.28	0.163	2.6	0.40
					Average =			2.8	0.43

Notes:
1. Tests were run at S&M (letter from Lois Baiad on 5/8/97). Equipment = FD2020 2000 watt power supply + CV15 xdu (ampl = 0.8 mils = 20.3 microns)
2. Lois Baiad reported (5/19/97) that the loss in air of the above stack with a green booster was 30 watts at full amplitude. I assume 20 watts for the converter (fixed) + 10 watts for the booster and horn.
3. For other boosters, I have assumed that the loss of the booster and horn (assumed 10 W on a green booster) increases with the square of the amplitude.
4. With a green booster @ 0.33 variac the approx amplitude is 0.33*20.3 = 6.7 microns. Since the power varies with the square of the ampl, the approx power is 0.33^2 * 30 watts = 3.3 watts.
	This compares to 2 watts for a 20 kHz 75mm Al spool horn with no booster (National Gypsum (99-031-01-00-0.dwg) (see Krell\20spool\3@5spl1\H22a))
4. Net cavitation power = Power in water - Power in air
5. For a horn whose face lateral dimensions are very large compared to the wavelength in water (approx. 7.4 cm @ 20 kHz), the power density may be higher than indicated above.

## Chart1a - S&M 20 kHz
Cavitation Power in Water at 20 kHz for Ø2.0" Unshaped Horn (Sonics & Materials)
Amplitude (microns, p-p)
Cavitation power density (watt/cm^2)
Amplitude (microns_peak)
Cavitation power density (watt/mm^2)

## Krell 20 kHz 3" spool data (1)
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$Krell 20 kHz 3" spool data (1)

20 kHz Ø3.0" Al spool horn for National Gypsum (99-031-01-00-0.dwg) (see Krell\20spool\3@5spl1\H22a)
S&M VCX-600 power supply + CV26 xdu (S/N 263262L) [my equipment]

	Face dia (inch)	Face area (in^2)	PS setting	Booster	Fotonic Sensor voltage (RMS)	Output ampl (microns)	Power in air (watts)	Power in water (watts)	Net cavitation power (watts)	Areal cavitation power density (watt/in^2)	Cavitation power density (watt/in^2/micron)	Cavitation power density (watt/cm^2/micron)
	3.000	7.069	100	None	0.360	26.4	18	461	443	63	2.4	0.37
			80	None	0.321	23.5	14	428	414	59	2.5	0.39
			60	None	0.266	19.5	10	333	323	46	2.3	0.36
			40	None	0.201	14.7	6	218	212	30	2.0	0.32
			30	None	0.161	11.8	2	153	151	21	1.8	0.28
			20	None	0.111	8.1	1	Wouldn't start	#VALUE!	#VALUE!
					Average =	2.2	0.36

	Important: Xdu hot and ground leads were probably reversed which gave higher than normal amplitudes
	Also, the amplitude was later discovered to have a "double" waveform on the scope and wouldn't synchronize.
	This problem was corrected when later running with the actual National Gypsum equipment (see tab "Krell 20 kHz 3" spool data (2)").

Notes:
1	All water power measurements use boiled (degassed) water
2	Water container = 7" dia x 3.5" deep plastic bowl
3	Horn immersed ~ 1/4" but doesn't make much difference as long as the water doesn't contact the undercut radius behind the flange
4	Horn output amplitude is calculated from the Fotonic Sensor calibration at 20 kHz (see "MTI-1000 calibration.xls", "Chart - Micro ampl vs Fr Slope", where slope is linear @ 73.2 microns p-p/v RMS up to 80 microns)
5	Power is measured with Clarke-Hess model 255 VAW meter, S/N 35424, calibrated at factory on 1/23/04. Wired per diagram at top of page following I-6.
6	Net cavitation power = Power in water - Power in air
7	Horn measured U = 98.8% (droop). See "20k 3.0 Al spool.xls"

## Krell 20 kHz 3" spool data (2)
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$Krell 20 kHz 3" spool data (2)
12/5/2005
20 kHz Ø3.0" Al spool horn for National Gypsum (99-031-01-00-0.dwg) (see Krell\20spool\3@5spl1\H22a)
S&M VC-750 power supply + CV33 xdu (S/N 3353381) [shipped equipment]

	Face dia (inch)	Face area (in^2)	PS setting	Booster	Fotonic Sensor voltage - test 1 (RMS)	Fotonic Sensor voltage - test 2 (RMS)	Fotonic Sensor voltage - average (RMS)	Output ampl (microns p-p)	Output ampl (microns peak)	Output velocity / (m/sec, peak)	Output velocity / (m/sec, RMS)	Power in air (watts)	Power in water (watts)	Total cavitation power (watts)	Estimated cavitation power due to radial side-amplitude (watts)	Net face cavitation power (watts)	Areal cavitation power density (watt/in^2)	Areal cavitation power density (watt/cm^2)	Areal cavitation power density (watt/mm^2)	Cavitation power density (watt/in^2/micron_pp)	Cavitation power density (watt/cm^2/micron_pp)		Areal cavitation power density (watt/m^2)	Areal radiation resistance (watt/cm^2)/(m/sec [rms])^2
	3.000	7.069	100	Gold	0.361	0.362	0.362	26.5	13.2	1.66	1.18	30	577	547	16	531	75	11.63	0.116	2.8	0.44		116341	8.42
			80	Gold	0.299	0.286	0.293	21.4	10.7	1.35	0.95	20	455	435	13	422	60	9.25	0.093	2.8	0.43		92520	10.23
			60	Gold	0.217	0.210	0.214	15.6	7.8	0.98	0.69	11	327	316	9	307	43	6.72	0.067	2.8	0.43		67210	13.94
			40	Gold	0.138	0.135	0.137	10.0	5.0	0.63	0.44	4	204	200	6	194	27	4.25	0.043	2.7	0.43		42538	21.59
			30	Gold	0.104	0.101	0.103	7.5	3.8	0.47	0.33	2	142	140	4	136	19	2.98	0.030	2.6	0.40		29776	26.80
			20	Gold	0.054	0.045	0.050	3.6	1.8	0.23	0.16	0	68	68	2	66	9	1.45	0.014	2.6	0.40		14463	55.82

	3.000	7.069	100	None	0.254	0.247	0.251	18.3	9.2	1.15	0.81	16	379	363	11	352	50	7.72	0.077	2.7	0.42		77206	11.64
			80	None	0.202	0.196	0.199	14.6	7.3	0.92	0.65	10	296	286	9	277	39	6.08	0.061	2.7	0.42		60829	14.53
			60	None	0.149	0.146	0.148	10.8	5.4	0.68	0.48	5	213	208	6	202	29	4.42	0.044	2.6	0.41		44239	19.23
			40	None	0.097	0.096	0.096	7.0	3.5	0.44	0.31	2	126	124	4	120	17	2.64	0.026	2.4	0.37		26373	26.92
			30	None	0.071	0.072	0.071	5.2	2.6	0.33	0.23	0	86	86	3	83	12	1.83	0.018	2.3	0.35		18291	33.93
			20	None	0.038	0.038	0.038	2.8	1.4	0.17	0.12	0	35	35	1	34	5	0.74	0.007	1.7	0.27		7444	48.76

Notes:
1	All water power measurements use boiled (degassed) water. Wavelength @ 20 kHz is approximatelly 2.92"
2	Water container = 7" dia x 3.5" deep plastic bowl (see below).
3	Horn immersed ~ 1/4" but doesn't make much difference as long as the water doesn't contact the undercut radius behind the flange
4	Horn output amplitude is calculated from the Fotonic Sensor calibration at 20 kHz (see "MTI-1000 calibration.xls", "Chart - Micro ampl vs Fr Slope", where slope is linear @ 73.2 microns p-p/v RMS up to 80 microns)
5	Power is measured with Clarke-Hess model 255 VAW meter, S/N 35424, calibrated at factory on 1/23/04. Wired per diagram at top of page following I-6.
6	Net cavitation power = Power in water - Power in air
7	Horn measured U = 98.8% (droop). See "20k 3.0 Al spool.xls"

	Plastic bowl used for cavitation tests:

## Chart2a - Krell 20k 3" spool
Cavitation Power in Water for 20 kHz Ø3.0" Al Spool Horn
Amplitude (microns, p-p)
Cavitation power density (watt/cm^2)
Amplitude (microns_peak)
Cavitation power density (watt/mm^2)

## Chart2a log - Krell 20k 3" spl
Krell Engineering Cavitation Power in Water for 20 kHz Ø3.0" Spool Horn
Amplitude (microns, p-p)
Cavitation power density (watt/cm^2)

## Chart2a log - Krell 20k 3"  (2)
Krell Engineering Cavitation Power in Water for 20 kHz Ø3.0" Spool Horn
RMS horn velocity (m/sec)
Cavitation power density (watt/m^2)

## Chart2a log - Krell 20k 3"  (3)
Krell Engineering Cavitation Power in Water for 20 kHz Ø3.0" Spool Horn
RMS horn face velocity (m/sec)
Cavitation power density (watt/m^2)

## Chart2b - Krell 20k 3" spool
Cavitation Power at 20 kHz
Amplitude (microns, p-p)
Cavitation power density (watt/cm^2/micron)

## Chart2c - Krell 20k 3" spool
Krell Engineering Cavitation Power in Water for 20 kHz Ø3.0" Spool Horn
Cavitation power density (watt/cm^2)
Areal radiation resistance (ohm/cm^2)

## 20 kHz cavitation horn data
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$20 kHz cavitation horn data
1-2-07
20 kHz Ø.626 Ti-6Al-4V cavitation horn for Naval Research (99-100-01-00-1.dwg) (see Krell\20cavitation\20cavitation2\H16a)
S&M VC-500 power supply (S/N 46042N) + CV33 xdu (S/N 3364114) + "Gold" (2.0 nominal) booster [shipped equipment]

					Power in water (watts)					Immersion depth = 0.24" (6.096 mm)					Immersion depth = 0.5" (12.7 mm)					Immersion depth = 1.0" (25.4 mm)
	Face dia (inch)	Face area (in^2)	PS setting	Booster	Fotonic Sensor voltage - test 1 (RMS)	Fotonic Sensor voltage - test 2 (RMS)	Fotonic Sensor voltage - average (RMS)	Output ampl (microns p-p)	Output ampl (microns peak)	Power in air (watts)	Immersion depth = .24"	Immersion depth = .5"	Immersion depth = 1.0"	Net cavitation power [Immersion depth = .25"] (watts)	Net cavitation power [Immersion depth = .5"] (watts)	Net cavitation power [Immersion depth = 1.0"] (watts)		Areal cavitation power density (watt/in^2)	Areal cavitation power density (watt/cm^2)	Areal cavitation power density (watt/mm^2)	Cavitation power density (watt/in^2/micron_pp)	Cavitation power density (watt/cm^2/micron_pp)		Areal cavitation power density (watt/cm^2)	Areal cavitation power density (watt/mm^2)	Cavitation power density (watt/in^2/micron_pp)	Cavitation power density (watt/cm^2/micron_pp)		Areal cavitation power density (watt/cm^2)	Areal cavitation power density (watt/mm^2)	Cavitation power density (watt/in^2/micron_pp)	Cavitation power density (watt/cm^2/micron_pp)
	0.626	0.308	100	Gold	0.972	0.972	0.9720	69.8	34.9	23	94	96	103	71	73	80		231	35.75	0.36	3.3	0.51		36.76	0.37	#NAME?	#NAME?		40.28	0.40	#NAME?	#NAME?
			90	Gold	0.885	0.880	0.8825	63.4	31.7	18	81	83	89	63	65	71		205	31.73	0.32	3.2	0.50		32.73	0.33	#NAME?	#NAME?		35.75	0.36	#NAME?	#NAME?
			80	Gold	0.787	0.785	0.7860	56.4	28.2	14	70	72	74	56	58	60		182	28.20	0.28	3.2	0.50		29.20	0.29	#NAME?	#NAME?		30.21	0.30	#NAME?	#NAME?
			60	Gold	0.597	0.582	0.5895	42.3	21.2	7	48	49	48	41	42	41		133	20.65	0.21	3.1	0.49		21.15	0.21	#NAME?	#NAME?		20.64	0.21	#NAME?	#NAME?
			40	Gold	0.409	0.400	0.4045	29.0	14.5	3	30	30	30	27	27	27		88	13.60	0.14	3.0	0.47		13.60	0.14	#NAME?	#NAME?		13.60	0.14	#NAME?	#NAME?
			20	Gold	0.171	0.170	0.1705	12.2	6.1	1	8	9	13	7	8	12		23	3.53	0.04	1.9	0.29		4.03	0.04	#NAME?	#NAME?		6.04	0.06	#NAME?	#NAME?

Notes:
1	All water power measurements use boiled (degassed) water. Wavelength @ 20 kHz is approximatelly 2.92". [Note: water may still have been somewhat warm from boiling, which could affect cavitation power. See Design.doc.]
2	Water container = 7" dia x 3.5" deep plastic bowl filled to lower rim (see below).
3	Horn output amplitude is calculated from the Fotonic Sensor calibration at 20 kHz (see see "MTI-1000 calibration.xls", "Chart - Front slope @ 20k", where slope is linear @ 71.8 microns p-p/v RMS up to ~70 microns)
4	Power is measured with Clarke-Hess model 255 VAW meter, S/N 35424, calibrated at factory on 1/23/04. Wired per diagram at top of page following I-6.
5	Net cavitation power = Power in water - Power in air
5	Transducer amplitude = 16.5 microns (measured at end of Solidica Ø1.5 M2 tool steel test horn)
6	"Gold" booster = S&M BHN11GD (measured G = 1.81 with Fotonic Sensor (12/26/06), 1.99 with Starrett 18-126 dial indicator)
6	Tip = Nitronic 60 (99-100-02-00-0.dwg)

	Plastic bowl used for cavitation tests:

## Chart3a - 20k cavitation horn
Cavitation Power in Water for 20 kHz Ø0.626" Cavitation Horn (ASTM G32) / (Immersion Depth = 0.24")
Amplitude (microns, p-p)
Cavitation power density (watt/cm^2)
Amplitude (microns_peak)
Cavitation power density (watt/mm^2)

## Chart3b - power vs immersion dp
Cavitation Power in Water for 20 kHz Ø0.626" Cavitation Horn (ASTM G32) (Immersion Depth = 0.24")
Amplitude (microns, p-p)
Cavitation power (watt)
Amplitude (microns, peak)
Cavitation power (watt)
Amplitude (microns_peak)
Cavitation power density (watt/mm^2)

## 20 kHz cav composite data (cm)
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$20 kHz cav composite data (cm)
11-10-2013

	Co (water, cm/sec) =	150000.0

					s			Peak-to-peak ampl				Peak ampl
	Horn	Frequency (kHz)	Water wavelength (cm)	Face dia (cm)	Face dia / Water wavelength	Face uniformity	Face area (cm^2)	Cavitation power density (watt/cm^2)/micron	Cavitation power density @ 20 kHz (watt/cm^2)/micron	Cavitation power density @ 20 kHz (watt/micron)		Cavitation power density (watt/cm^2)/micron	Cavitation power density @ 20 kHz (watt/cm^2)/micron	Cavitation power density @ 20 kHz (watt/micron)	Source data
	3.0" dia spool	20.0	7.500	7.620	1.016	0.995	45.604	0.453	0.453	20.658		0.906	0.906	41.317	Don Culp - slope of graph "Chart2a - Krell 20k 3" spool"
	60 mm dia	17.8	8.427	6.000	0.712		28.274	0.395	0.444	12.549		0.790	0.888	25.097	Peshkovsky - "Matching a transducer to water at cavitation: Acoustic horn design principles", Ultrasonics Sonochemistry 14 (2007), pp. 314 - 322 (see fig 9, p. 321); / Slope = (1e5 W/m^2)/(m_rms/sec) / = (0.707 W/m^2)/(m_peak/sec) / = (0.707e-5 w/cm^2)/(micron_peak/sec)
	2.0" dia (S&M data)	20.0	7.500	5.080	0.677		20.268	0.385	0.385	7.803		0.770	0.770	15.607	Sonics & Materials - slope of graph "Chart - S&M 20 kHz"
	19 mm dia	20.0	7.500	1.900	0.253	~1.00	2.835	0.530	0.530	1.503		1.060	1.060	3.005	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")
	0.626" dia cavitation test horn	20.0	7.500	1.590	0.212	~1.00	1.986	0.552	0.552	1.096		1.104	1.104	2.192	Don Culp - slope of graph "Chart3a - 20k cavitation horn"
	9.5 mm dia	20.0	7.500	0.950	0.127	~1.00	0.709	0.549	0.549	0.389		1.098	1.098	0.778	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")
	3 mm dia	20.0	7.500	0.300	0.040	~1.00	0.071	1.410	1.410	0.100		2.820	2.820	0.199	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")

Notes:
1	Power density at 20 kHz assumes that, for a given amplitude, the power varies directly with the frequency.

	2.0" dia (S&M data)	20.0	7.500	5.080	0.677		20.268	0.385	0.385	7.803		0.770	0.770	15.607	Sonics & Materials - slope of graph "Chart - S&M 20 kHz"

## Chart4a - cav composite data (1)
Horn face diameter (cm)
Cavitation power intensity (watt/cm^2/micron p-p)

## Chart4a - cav composite data (2)
Horn face diameter (cm)
Cavitation power intensity (watt/cm^2/micron p-p)

## Chart4a - cav composite data (3)
Horn face diameter (cm)
Cavitation power intensity (watt/cm^2/micron p-p)

## Chart4a - cav composite data (3a)
Horn face area (cm^2)
Cavitation power intensity ((watt/micron p-p)/cm^2)
Horn diameter / Water wavelength
Cavitation power intensity ((watt/micron p-p)/cm^2)

## Chart4a - cav composite data (4)
Ampl – microns peak
Horn face area (cm^2)
Cavitation power (watt/micron p-p)
Horn face area (cm^2)
Cavitation power (watt/micron_peak)

## 20 kHz cav composite data (mm)
'file:///E:/websites/usr/miscellaneous/wear/cavitation/_images/_support_files/Cavitation_1k.ods'#$20 kHz cav composite data (mm)
11-10-2013

	Co (water, m/sec) =	1500.0

					Peak-to-peak ampl				Peak ampl
	Horn	Frequency (kHz)	Water wavelength (mm)	Face dia (mm)	Face dia / Water wavelength	Face uniformity	Face area (mm^2)	Cavitation power density (watt/mm^2)/micron	Cavitation power density @ 20 kHz (watt/mm^2)/micron	Cavitation power density @ 20 kHz (watt/micron)		Cavitation power density (watt/mm^2)/micron	Cavitation power density @ 20 kHz (watt/mm^2)/micron	Cavitation power density @ 20 kHz (watt/micron)	Source data
	3.0" dia spool	20.0	75.0	76.200	1.016	0.995	4560.4	0.00453	0.00453	20.658		0.00906	0.00906	41.317	Don Culp - slope of graph "Chart2a - Krell 20k 3" spool"
	60 mm dia	17.8	84.3	60.000	0.712		2827.4	0.00395	0.00444	12.549		0.00790	0.00888	25.097	Peshkovsky - "Matching a transducer to water at cavitation: Acoustic horn design principles", Ultrasonics Sonochemistry 14 (2007), pp. 314 - 322 (see fig 9, p. 321); / Slope = (1e5 W/m^2)/(m_rms/sec) / = (0.707 W/m^2)/(m_peak/sec) / = (0.707e-5 w/cm^2)/(micron_peak/sec)
	2.0" dia (S&M data)	20.0	75.0	50.800	0.677		2026.8	0.00385	0.00385	7.803		0.00770	0.00770	15.607	Sonics & Materials - slope of graph "Chart - S&M 20 kHz"
	19 mm dia	20.0	75.0	19.000	0.253	~1.00	283.5	0.00530	0.00530	1.503		0.01060	0.01060	3.005	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")
	0.626" dia cavitation test horn	20.0	75.0	15.900	0.212	~1.00	198.6	0.00552	0.00552	1.096		0.01104	0.01104	2.192	Don Culp - slope of graph "Chart3a - 20k cavitation horn"
	9.5 mm dia	20.0	75.0	9.500	0.127	~1.00	70.9	0.00549	0.00549	0.389		0.01098	0.01098	0.778	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")
	3 mm dia	20.0	75.0	3.000	0.040	~1.00	7.1	0.01410	0.01410	0.100		0.02820	0.02820	0.199	Write\Ultrasonics\Cavitation\”Perkins, 'Power Ultrasonic Equipment - Practice and Application'.pdf” / from Sonic Systems (http://sonicsystems.co.uk/page/power-ultrasonics-a-guide/39/) / (Also see \Write\_Design\Design.doc, section "Cavitation")

Notes:
1	Power density at 20 kHz assumes that, for a given amplitude, the power varies directly with the frequency.

	2.0" dia (S&M data)	20.0	0.1	5.080	67.733		20.3	0.38500	0.38500	7.803		0.77000	0.77000	15.607	Sonics & Materials - slope of graph "Chart - S&M 20 kHz"

## Chart5a - cav composite data (3a)
Horn face area (mm^2)
Cavitation power intensity ((watt/micron p-p)/mm^2)
Horn diameter / Water wavelength
Cavitation power intensity ((watt/micron p-p)/mm^2)
Horn face area (mm^2)
Cavitation power intensity ((watt/micron_peak)/mm^2)
Horn diameter / Water wavelength
Cavitation power intensity ((watt/micron_peak)/mm^2)

## Chart5a - cav composite data (4)
Ampl – microns peak
Horn face area (mm^2)
Cavitation power (watt/micron p-p)
Horn face area (mm^2)
Cavitation power (watt/micron_peak)